## PBNP
Patientennr. | Vista (pg/ml) | Finecare (ng/L) |  |  | Vista (pg/ml) | Finecare (ng/L) | Bland-Altman-plot | Finecare-Vista (ng/ml) | Mean Finecare_Vista | Finecare/Vista
22648202 | 7392.7 | 4432 | Performance: |  |  |  |  | -2960.7 | 5912.35 | 0.59951
22648273 | 1252.3 | 546.3 | [5 empty] | -706 | 899.3 | 0.43624
22648637 | 1888.8 | 888.4 | Analyrical measurement range: |  | 5 - 35000 | 18 - 35000 |  | -1000.4 | 1388.6 | 0.47035
22649528 | 7634.5 | 7643.8 | [5 empty] | 9.3 | 7639.15 | 1.0012
22650186 | 4759 | 2971.9 | Expected values: | < 75 years | <125 | <300 |  | -1787.1 | 3865.45 | 0.62448
22649687 | 2471.3 | 631 |  | > 75 years | <460 | <450 |  | -1840.3 | 1551.15 | 0.25533
22650604 | 1991.3 | 1481.7 | [5 empty] | -509.6 | 1736.5 | 0.74409
22650844 | 1353.2 | 304.2 | [5 empty] | -1049 | 828.7 | 0.2248
22651072 | 12211 | 6120.1 | [5 empty] | -6090.9 | 9165.55 | 0.5012
22652131 | 564.5 | 351.7 | [5 empty] | -212.8 | 458.1 | 0.62303
22653543 | 349 | 229 | [5 empty] | -120 | 289 | 0.65616
22653274 | 767.3 | 509.5 | [5 empty] | -257.8 | 638.4 | 0.66402
22653484 | 1817.9 | 318.5 | [5 empty] | -1499.4 | 1068.2 | 0.1752
22655839 | 476.8 | 199.4 | [5 empty] | -277.4 | 338.1 | 0.4182
22656560 | 600.2 | 264 | [5 empty] | -336.2 | 432.1 | 0.43985
22660117 | 8923 | 3958 | [5 empty] | -4965 | 6440.5 | 0.44357
22660582 | 971 | 389 | [5 empty] | -582 | 680 | 0.40062
22661862 | 9862.9 | 4474.2 | [5 empty] | -5388.7 | 7168.55 | 0.45364
22663279 | 4963.8 | 1439.4 | [5 empty] | -3524.4 | 3201.6 | 0.28998
22663282 | 3929.1 | 2031.5 | [5 empty] | -1897.6 | 2980.3 | 0.51704
22663030 | 151.2 | 122 | [5 empty] | -29.2 | 136.6 | 0.80688
22662906 | 78.3 | 36.5 | [5 empty] | -41.8 | 57.4 | 0.46616
22663376 | 5630.1 | 1718.2 | [5 empty] | -3911.9 | 3674.15 | 0.30518
22664297 | 899.5 | 559.2 | [5 empty] | -340.3 | 729.35 | 0.62168
22663369 | 12392 | 7610 | [5 empty] | -4782 | 10001 | 0.61411
22664802 | 1120.3 | 340.4 | [5 empty] | -779.9 | 730.35 | 0.30385
22664895 | 279 | 236.7 | [5 empty] | -42.3 | 257.85 | 0.84839
22664836 | 1949.1 | 1148.6 | [5 empty] | -800.5 | 1548.85 | 0.5893
22664911 | 317.5 | 158.6 | [5 empty] | -158.9 | 238.05 | 0.49953
[7 empty] | MW difference | -1582.1655 | 2553.6276 | 0.51702
[7 empty] | Stdabw | 1840.9051 | 2950.3757 | 0.19264
Patientennr. | Vista (pg/ml) | Finecare (ng/L) |  |  |  |  | Limit -2s | -5263.9757 |  | 0.13174
22662906 | 78.3 | 36.5 |  |  |  |  | Limit +2s | 2099.6447 |  | 0.9023
22663030 | 151.2 | 122
22664895 | 279 | 236.7 |  |  |  |  | Bland-Altman-plot | Finecare-Vista (ng/ml)
22664911 | 317.5 | 158.6 | [5 empty] | -41.8 | 57.4
22653543 | 349 | 229 | [5 empty] | -29.2 | 136.6
22655839 | 476.8 | 199.4 | [5 empty] | -42.3 | 257.85
22652131 | 564.5 | 351.7 | [5 empty] | -158.9 | 238.05
22656560 | 600.2 | 264 | [5 empty] | -120 | 289
22653274 | 767.3 | 509.5 | [5 empty] | -277.4 | 338.1
22664297 | 899.5 | 559.2 | [5 empty] | -212.8 | 458.1
22660582 | 971 | 389 | [5 empty] | -336.2 | 432.1
22664802 | 1120.3 | 340.4 | [5 empty] | -257.8 | 638.4
22648273 | 1252.3 | 546.3 | [5 empty] | -340.3 | 729.35
22650844 | 1353.2 | 304.2 | [5 empty] | -582 | 680
22653484 | 1817.9 | 318.5 | [5 empty] | -779.9 | 730.35
22648637 | 1888.8 | 888.4 | [5 empty] | -706 | 899.3
22664836 | 1949.1 | 1148.6 | [5 empty] | -1049 | 828.7
22650604 | 1991.3 | 1481.7 | [5 empty] | -1499.4 | 1068.2
[8 empty] | -1000.4 | 1388.6
[8 empty] | -800.5 | 1548.85
[8 empty] | -509.6 | 1736.5
[7 empty] | MW d | -485.75
[7 empty] | Stdabw | 413.7142
[7 empty] | Grenzen 2s | -1313.1785
Comment:
The "Bland-Altmann" Finecare/Vista ratio shows, that the values of Finecare over the whole measurement range are 50% lower in average and there is a big variation in the low measurement range.
Comment:
The difference in the mean is -485,74 ng/ml Finecare to Vista in the range 0-2000 pg/ml.
Further the "Bland-Altmann" shows, that there is a higher varability of differences between Fincare and Vista in dependence to higher meaurements results!
Comment: There is a moderate correlation over the whole measurement range. If we look at the lower range near the expected value, the correlation coefficient is less.
The mean difference is -1500 ng/L Finecare to Vista of whole range.

## CRP
Patientennr. | Vista | Finecare | hcCRP | Vista 10*mg/dl | Finecare |  | Vista | Finecare |  | Finecare-Vista (ng/ml) | Mean Finecare Vista | Finecare/Vista
22646677 | 17 | 120.8 |  | 170 | 120.8 | Performance: |  | It was difficult to puncture the buffer tube with the pipette with the sample in. Then there first puncture with empty pipette and after that the sample |  | -49.2 | 145.4 | #DIV/0!
22648104 | <0,29 | <5 | 1.3 | 2.9 | 1.3 |  |  |  |  | -1.6 | 2.1
22648042 | 5.17 | 43.4 | >5 | 51.7 | 43.4 |  |  | That means the red marked results are not realted to that. |  | -8.3 | 47.55
22648099 | 0.47 | <5 | 4.7 | 4.7 | 4.7 | Analytical Measurement Range: | 2,9-190 mg/l | 0,5-200 mg/l |  | 0 | 4.7
22648080 | 1.35 | 14.1 | >5 | 13.5 | 14.1 | Questions | Results above 1,9 mg/L have been measured again in diluted form? |  |  | 0.6 | 13.8
22648116 | 1.03 | 10.4 | >5 | 10.3 | 10.4 |  |  |  |  | 0.1 | 10.35
22647868 | 20.7 | >200 | >5 | 207 | 200 |  |  |  |  | -7 | 203.5
22647853 | 12 | 115.5 | >5 | 120 | 115.5 |  |  |  |  | -4.5 | 117.75
22647869 | 0.65 | 7.3 | >5 | 6.5 | 7.3 |  |  |  |  | 0.8 | 6.9
22648002 | 6.09 | 50.9 | >5 | 60.9 | 50.9 |  |  |  |  | -10 | 55.9
22648073 | 4.15 | 40.4 | >5 | 41.5 | 40.4 |  |  |  |  | -1.1 | 40.95
22648030 | 0.33 | <5 | 2.4 | 3.3 | 2.4 |  |  |  |  | -0.9 | 2.85
22647944 | 9.08 | 99.6 | >5 | 90.8 | 99.6 |  |  |  |  | 8.8 | 95.2
22647573 | 2.14 | 23.4 | >5 | 21.4 | 23.4 |  |  |  |  | 2 | 22.4
22647857 | 0.39 | <5 | 3.5 | 3.9 | 3.5 |  |  |  |  | -0.4 | 3.7
22649222 | 25 | 188.6 | >5 | 250 | 188.6 |  |  |  |  | -61.4 | 219.3
22649250 | 0.72 | 60.3 | >5 | 7.2 | 60.3 |  |  |  |  | 53.1 | 33.75
22649556 | 3.74 | 27.3 | >5 | 37.4 | 27.3 |  |  |  |  | -10.1 | 32.35
22649584 | 0.7 | 6.4 | >5 | 7 | 6.4 |  |  |  |  | -0.6 | 6.7
22649566 | <0,29 | <5 | <0,5 | 2.9 | 0.5 |  |  |  |  | -2.4 | 1.7
22649568 | 1.55 | 14.8 | >5 | 15.5 | 14.8 |  |  |  |  | -0.7 | 15.15
22649579 | 18.8 | 127.5 | >5 | 188 | 127.5 |  |  |  |  | -60.5 | 157.75
22649531 | 20.3 | 115.7 | >5 | 203 | 115.7 |  |  |  |  | -87.3 | 159.35
22649609 | 4.05 | 30.2 | >5 | 40.5 | 30.2 |  |  |  |  | -10.3 | 35.35
22649618 | 0.58 | 5 | 5 | 5.8 | 5 |  |  |  |  | -0.8 | 5.4
22649624 | 1.03 | 10.8 | >5 | 10.3 | 10.8 |  |  |  |  | 0.5 | 10.55
22649479 | <0,29 | <5 | 1 | 2.9 | 1 |  |  |  |  | -1.9 | 1.95
22649210 | <0,29 | <5 | <0,5 | 2.9 | 0.5 |  |  |  |  | -2.4 | 1.7
22651363 | 7.33 | 128.2 | >5 | 73.3 | 128.2 |  |  |  |  | 54.9 | 100.75
22651080 | 18.4 | 111.3 | >5 | 184 | 111.3 |  |  |  |  | -72.7 | 147.65
22651397 | 22.5 | 183.3 | >5 | 225 | 183.3 |  |  |  |  | -41.7 | 204.15
22653134 | 15.8 | 110.9 | >5 | 158 | 110.9 |  |  |  |  | -47.1 | 134.45
22652840 | 0.72 | 8.5 | >5 | 7.2 | 8.5 |  |  |  |  | 1.3 | 7.85
22652852 | 42.3 | 194.1 | >5 | 423 | 194.1 |  |  |  |  | -228.9 | 308.55
22652608 | <0,29 | <0,5 | 1.6 | 2.9 | 1.6 |  |  |  |  | -1.3 | 2.25
22652587 | 1.24 | 13.1 | >5 | 12.4 | 13.1 |  |  |  |  | 0.7 | 12.75
22655854 | 19.7 | 161.3 | >5 | 197 | 161.3 |  |  |  |  | -35.7 | 179.15
22651385 | 0.94 | 7.4 | >5 | 9.4 | 7.4 |  |  |  |  | -2 | 8.4
22656576 | <0,29 | <0,5 | 0.5 | 2.9 | 0.5 |  |  |  |  | -2.4 | 1.7
22656573 | 10 | 64.9 | >5 | 100 | 64.9 |  |  |  |  | -35.1 | 82.45
22656581 | 14.4 | 151.2 | >5 | 144 | 151.2 |  |  |  |  | 7.2 | 147.6
22656708 | 0.45 | 4.1 | >5 | 4.5 | 4.1 |  |  |  |  | -0.4 | 4.3
22656410 | 9.35 | 21.5 | >5 | 93.5 | 21.5 |  |  |  |  | -72 | 57.5
22656829 | 11.5 | 109.8 | >5 | 115 | 109.8 |  |  |  |  | -5.2 | 112.4
22656832 | 20 | >200 | >5 | 200 | 200 |  |  |  |  | 0 | 200
22657168 | 2.22 | 16.3 | >5 | 22.2 | 16.3 |  |  |  |  | -5.9 | 19.25
22657116 | 9.56 | 70.2 | >5 | 95.6 | 70.2 |  |  |  |  | -25.4 | 82.9
22657082 | 22.3 | 145.7 | >5 | 223 | 145.7 |  |  |  |  | -77.3 | 184.35
22656929 | 3.3 | 22.3 | >5 | 33 | 22.3 |  |  |  |  | -10.7 | 27.65
22660160 | 31.4 | 51.6 | >5 | 314 | 51.6 |  |  |  |  | -262.4 | 182.8
22660319 | 29.8 | 59.3 | >5 | 298 | 59.3 |  |  |  |  | -238.7 | 178.65
22660248 | 42.5 | 134.9 | >5 | 425 | 134.9 |  |  |  |  | -290.1 | 279.95
22660422 | <0,29 | <5 | 3.1 | 2.9 | 3.1 |  |  |  |  | 0.2 | 3
22660311 | 12.5 | 84.5 | >5 | 125 | 84.5 |  |  |  |  | -40.5 | 104.75
22643277 | 1.14 | 7.9 | >5 | 11.4 | 7.9 |  |  |  |  | -3.5 | 9.65
22641093 | 10 | 84 | >5 | 100 | 84 |  |  |  |  | -16 | 92
22641092 | 23 | 122.5 | >5 | 230 | 122.5 |  |  |  |  | -107.5 | 176.25
22641094 | 0.62 | 7 | >5 | 6.2 | 7 |  |  |  |  | 0.8 | 6.6
[9 empty] | MW difference | -31.2569 | 77.8233
[9 empty] | StdAbw | 68.4203 | 82.1407
[9 empty] | Limit -2s | -168.0974
[9 empty] | Limit +2s | 105.5837
[10 empty] | Finecare-Vista (ng/ml) | Mean Finecare Vista | Finecare/Vista
|  |  |  | Vista 10*mg/dl | Finecare |  |  |  |  | -1.6 | 2.1 | #VALUE!
|  |  |  | 2.9 | 1.3 |  |  |  |  | -2.4 | 1.7
|  |  |  | 2.9 | 0.5 |  |  |  |  | -1.9 | 1.95
|  |  |  | 2.9 | 1 |  |  |  |  | -2.4 | 1.7
|  |  |  | 2.9 | 0.5 |  |  |  |  | -1.3 | 2.25
|  |  |  | 2.9 | 1.6 |  |  |  |  | -2.4 | 1.7
|  |  |  | 2.9 | 0.5 |  |  |  |  | 0.2 | 3
|  |  |  | 2.9 | 3.1 |  |  |  |  | -0.9 | 2.85
|  |  |  | 3.3 | 2.4 |  |  |  |  | -0.4 | 3.7
|  |  |  | 3.9 | 3.5 |  |  |  |  | -0.4 | 4.3
|  |  |  | 4.5 | 4.1 |  |  |  |  | 0 | 4.7
|  |  |  | 4.7 | 4.7 |  |  |  |  | -0.8 | 5.4
|  |  |  | 5.8 | 5 |  |  |  |  | 0.8 | 6.6
|  |  |  | 6.2 | 7 |  |  |  |  | 0.8 | 6.9
|  |  |  | 6.5 | 7.3 |  |  |  |  | -0.6 | 6.7
|  |  |  | 7 | 6.4 |  |  |  |  | 53.1 | 33.75
|  |  |  | 7.2 | 60.3 |  |  |  |  | 1.3 | 7.85
|  |  |  | 7.2 | 8.5 |  |  |  |  | -2 | 8.4
|  |  |  | 9.4 | 7.4 |  |  |  |  | 0.1 | 10.35
|  |  |  | 10.3 | 10.4 |  |  |  |  | 0.5 | 10.55
|  |  |  | 10.3 | 10.8 |  |  |  |  | -3.5 | 9.65
|  |  |  | 11.4 | 7.9 |  |  |  |  | 0.7 | 12.75
|  |  |  | 12.4 | 13.1 |  |  |  |  | 0.6 | 13.8
|  |  |  | 13.5 | 14.1 |  |  |  |  | -0.7 | 15.15
|  |  |  | 15.5 | 14.8 |  |  |  |  | 2 | 22.4
|  |  |  | 21.4 | 23.4 |  |  |  |  | -5.9 | 19.25
|  |  |  | 22.2 | 16.3 |  |  |  |  | -10.7 | 27.65
|  |  |  | 33 | 22.3 |  |  |  |  | -10.1 | 32.35
|  |  |  | 37.4 | 27.3 |  |  |  |  | -10.3 | 35.35
|  |  |  | 40.5 | 30.2 |  |  |  |  | -1.1 | 40.95
|  |  |  | 41.5 | 40.4 |  |  |  |  | -8.3 | 47.55
|  |  |  | 51.7 | 43.4 |  |  |  |  | -10 | 55.9
|  |  |  | 60.9 | 50.9 |  |  |  |  | 54.9 | 100.75
|  |  |  | 73.3 | 128.2 |  |  |  |  | 8.8 | 95.2
|  |  |  | 90.8 | 99.6 |  |  |  |  | -72 | 57.5
|  |  |  | 93.5 | 21.5 |  |  |  |  | -25.4 | 82.9
|  |  |  | 95.6 | 70.2 |  |  |  |  | -35.1 | 82.45
|  |  |  | 100 | 64.9 |  |  |  |  | -16 | 92
|  |  |  | 100 | 84 |  |  |  |  | -5.2 | 112.4
|  |  |  | 115 | 109.8 |  |  |  |  | -4.5 | 117.75
|  |  |  | 120 | 115.5 |  |  |  |  | -40.5 | 104.75
|  |  |  | 125 | 84.5 |  |  |  |  | 7.2 | 147.6
|  |  |  | 144 | 151.2 |  |  |  |  | -47.1 | 134.45
|  |  |  | 158 | 110.9 |  |  |  |  | -49.2 | 145.4
|  |  |  | 170 | 120.8 |  |  |  |  | -72.7 | 147.65
|  |  |  | 184 | 111.3 |  |  |  |  | -60.5 | 157.75
|  |  |  | 188 | 127.5 |  |  |  |  | -35.7 | 179.15
|  |  |  | 197 | 161.3 |  |  |  |  | 0 | 200
|  |  |  | 200 | 200 |  |  |  | MW difference | -10.0384 | 52.9882
[9 empty] | StdAbw | 38.5301 | 57.6553
[9 empty] | Limit -2s | -87.0986
[9 empty] | Limit +2s | 67.0218
Comment: In the insert sheet of the reference method Vista is written, that the analytical measurement range is 2.9 to 190 mg/L.
Sample that show results above 190 mg/L has to be diluted.
So I looked for the correlation coefficient in this range.
Result: The correlation is better in the range from 0-200 mg/L.
Comment: Red marked are Vista results above 1,9 mg/L
The mean difference is -31,26 ng/ml Finecare to Vista of whole range.
The difference in the mean is -10 ng/ml Finecare to Vista in the range 0-200 ng/ml.
Further the "Bland-Altmann" shows, that there is a higher varability of differences between Fincare and Vista in dependence to higher meaurements results!

## HbA1c
Patientennr. | Vista | Finecare |  | Vista | Finecare
22648083 | 5.3 | 5.5
22647881 | 5.5 | 4.2 | Durchführung:
22648423 | 5.2 | 4.9
22648314 | 9.1 | 8.6 | AMR: | 3,5-16 % | 4-14,5 %
22648357 | 6.3 | 5.7 | erwartete Werte: | 4,2-6,3 %
22648308 | 5.4 | 5.1
22648481 | 5.8 | 5.3 | Bemerkung: | gute Korrelation
22648794 | 6.3 | 5.8
22648692 | 5.9 | 5.2
22649017 | 6.7 | 6.1
22649051 | 9.1 | 8.4
22649412 | 6.3 | 6
22649229 | 5.3 | 4.9
22648817 | 7.8 | 7.5
22650491 | 4.9 | 4.5
22649777 | 6.6 | 5.9
22650912 | 4.8 | 4
22650828 | 5.4 | 5
22650803 | 5.5 | 5
22650968 | 9.7 | 9.4
22650929 | 5.1 | 4.8
22651510 | 5 | 5.2
22652097 | 7.1 | 6.8
22653180 | 6.2 | 6.5
22653497 | 5.1 | 5.3
22653262 | 5.6 | 5.9
22653448 | 7.1 | 6.6
22650449 | 7.9 | 8
22653259 | 6.4 | 6.1
22653939 | 5 | 5.4
22653891 | 8.6 | 8.7
22653944 | 7.3 | 6.9
50023943 | 6.7 | 6.6
50023944 | 4.9 | 4.7
22654377 | 5.3 | 5.3
22654758 | 9 | 8.2
22654808 | 5.7 | 6.6
50023967 | 12.7 | 11.9
22654769 | 8.1 | 8.4
50023991 | 7.6 | 7.8
22656402 | 5.4 | 5.4
22656725 | 7.9 | 8.3
50025180 | 6.1 | 6.1
22659720 | 6.5 | 6.1
Comment: HbA1c shows the best correlation to reference method.

## D-Dim
Patientennr. | Vidas | Finecare |  |  | Vista ng/ml | Finecare mg/L |  | Finecare-Vista (ng/ml) | Mean Finecare Vista
22648202 | 1.02 | 1.5 | [5 empty] | 0.48 | 1.26
22648210 | 0.87 | 1.2 | [5 empty] | 0.33 | 1.035
22648600 | 0.35 | 0.5 | AMR |  | 45-10000 | 0,1-10 |  | 0.15 | 0.425
22648702 | 1.02 | 2.9 | [5 empty] | 1.88 | 1.96
22648658 | 6.31 | 10 | clinical decision limit: | venous thrmboembolism | 500 | 0.5 |  | 3.69 | 8.155
22648987 | 1.97 | 6.1 | [5 empty] | 4.13 | 4.035
22649000 | 6.71 | 10 | [5 empty] | 3.29 | 8.355
22649003 | 0.24 | 0.5 | [5 empty] | 0.26 | 0.37
22649236 | 0.27 | 0.6 | [5 empty] | 0.33 | 0.435
22649481 | 0.23 | 0.4 | [5 empty] | 0.17 | 0.315
22649567 | 0.98 | 0.8 | [5 empty] | -0.18 | 0.89
22649522 | 0.94 | 1.5 | [5 empty] | 0.56 | 1.22
22649551 | 1.33 | 3 | [5 empty] | 1.67 | 2.165
22650464 | 3.66 | 3.6 | [5 empty] | -0.06 | 3.63
22650515 | 0.41 | 0.7 | [5 empty] | 0.29 | 0.555
22650775 | 0.38 | 0.6 | [5 empty] | 0.22 | 0.49
22651397 | 0.72 | 1.4 | [5 empty] | 0.68 | 1.06
22651437 | 0.89 | 1 | [5 empty] | 0.11 | 0.945
22652125 | 0.39 | 0.6 | [5 empty] | 0.21 | 0.495
22651643 | 0.19 | 0.5 | [5 empty] | 0.31 | 0.345
22652114 | 8.38 | 10 | [5 empty] | 1.62 | 9.19
22653240 | 0.17 | 0.4 | [5 empty] | 0.23 | 0.285
22653201 | 1.78 | 2.4 | [5 empty] | 0.62 | 2.09
22653313 | 3.84 | 9 | [5 empty] | 5.16 | 6.42
22653614 | 0.39 | 5.4 | [5 empty] | 5.01 | 2.895
22653652 | 0.45 | 1 | [5 empty] | 0.55 | 0.725
22656783 | 0.33 | 1 | [5 empty] | 0.67 | 0.665
22655893 | 0.28 | 0.8 | [5 empty] | 0.52 | 0.54
22656154 | 4.64 | 10 | [5 empty] | 5.36 | 7.32
22656321 | 1.8 | 3.4 | [5 empty] | 1.6 | 2.6
22655934 | 2.39 | 5.7 | [5 empty] | 3.31 | 4.045
22656366 | 0.16 | 0.6 | [5 empty] | 0.44 | 0.38
22652494 | 0.22 | 0.6 | [5 empty] | 0.38 | 0.41
22656492 | 0.33 | 0.9 | [5 empty] | 0.57 | 0.615
22658634 | 0.37 | 0.7 | [5 empty] | 0.33 | 0.535
22658966 | 0.81 | 1.9 | [5 empty] | 1.09 | 1.355
22652358 | 0.2 | 0.5 | [5 empty] | 0.3 | 0.35
22654827 | 0.34 | 0.5 | [5 empty] | 0.16 | 0.42
22654794 | 0.59 | 1.9 | [5 empty] | 1.31 | 1.245
22655265 | 1 | 2.4 | [5 empty] | 1.4 | 1.7
22655291 | 0.17 | 0.3 | [5 empty] | 0.13 | 0.235
22655255 | 1.23 | 2.4 | [5 empty] | 1.17 | 1.815
22658480 | 0.22 | 0.66 | [5 empty] | 0.44 | 0.44
22658750 | 2.79 | 4.9 | [5 empty] | 2.11 | 3.845
22643361 | 0.33 | 0.8 | [5 empty] | 0.47 | 0.565
[7 empty] | MW difference | 1.1882 | 1.9739
[7 empty] | StdAbw | 1.4772 | 2.3842
Patientennr. | Vidas | Finecare |  |  |  |  | Limit -2s | -1.7661
22656366 | 0.16 | 0.6 |  |  |  |  | Limit +2s | 4.1426
22653240 | 0.17 | 0.4 | [5 empty] | Vidas | Finecare
22655291 | 0.17 | 0.3 | [5 empty] | 0.44 | 0.38
22651643 | 0.19 | 0.5 | [5 empty] | 0.23 | 0.285
22652358 | 0.2 | 0.5 | [5 empty] | 0.13 | 0.235
22652494 | 0.22 | 0.6 | [5 empty] | 0.31 | 0.345
22658480 | 0.22 | 0.66 | [5 empty] | 0.3 | 0.35
22649481 | 0.23 | 0.4 | [5 empty] | 0.38 | 0.41
22649003 | 0.24 | 0.5 | [5 empty] | 0.44 | 0.44
22649236 | 0.27 | 0.6 | [5 empty] | 0.17 | 0.315
22655893 | 0.28 | 0.8 | [5 empty] | 0.26 | 0.37
22656783 | 0.33 | 1 | [5 empty] | 0.33 | 0.435
22656492 | 0.33 | 0.9 | [5 empty] | 0.52 | 0.54
22643361 | 0.33 | 0.8 | [5 empty] | 0.67 | 0.665
22654827 | 0.34 | 0.5 | [5 empty] | 0.57 | 0.615
22648600 | 0.35 | 0.5 | [5 empty] | 0.47 | 0.565
22658634 | 0.37 | 0.7 | [5 empty] | 0.16 | 0.42
22650775 | 0.38 | 0.6 | [5 empty] | 0.15 | 0.425
22652125 | 0.39 | 0.6 | [5 empty] | 0.33 | 0.535
22653614 | 0.39 | 5.4 | [5 empty] | 0.22 | 0.49
22650515 | 0.41 | 0.7 | [5 empty] | 0.21 | 0.495
22653652 | 0.45 | 1 | [5 empty] | 5.01 | 2.895
22654794 | 0.59 | 1.9 | [5 empty] | 0.29 | 0.555
22651397 | 0.72 | 1.4 | [5 empty] | 0.55 | 0.725
22658966 | 0.81 | 1.9 | [5 empty] | 1.31 | 1.245
22648210 | 0.87 | 1.2 | [5 empty] | 0.68 | 1.06
22651437 | 0.89 | 1 | [5 empty] | 1.09 | 1.355
22649522 | 0.94 | 1.5 | [5 empty] | 0.33 | 1.035
22649567 | 0.98 | 0.8 | [5 empty] | 0.11 | 0.945
22655265 | 1 | 2.4 | [5 empty] | 0.56 | 1.22
22648202 | 1.02 | 1.5 | [5 empty] | -0.18 | 0.89
22648702 | 1.02 | 2.9 | [5 empty] | 1.4 | 1.7
22655255 | 1.23 | 2.4 | [5 empty] | 0.48 | 1.26
22649551 | 1.33 | 3 | [5 empty] | 1.88 | 1.96
22653201 | 1.78 | 2.4 | [5 empty] | 1.17 | 1.815
22656321 | 1.8 | 3.4 | [5 empty] | 1.67 | 2.165
22648987 | 1.97 | 6.1 | [5 empty] | 0.62 | 2.09
[8 empty] | 1.6 | 2.6
[8 empty] | 4.13 | 4.035
[7 empty] | MW difference | 0.78351 | 1.0234
[7 empty] | StdAbw | 1.0408 | 0.86278
[7 empty] | Limit -2s | -1.2981
[7 empty] | Limit +2s | 2.8651
Comment: If we look at the relavant range for clinical diagnose, so the correlation between Finecar and Vidas is less, than over the whole range.
It seems that Finecare cut-off is rather 1 mg/L than 0,5 mg/L, because the patients show the same clinical appearance.
The mean difference is 1,19 mg/L Finecare to Vista of whole range.
The difference in the mean is 0,78 mg/L Finecare to Vista in the range 0-2 mg/L.
Further the "Bland-Altmann" shows, that there is a higher varability of differences between Fincare and Vista in dependence to higher meaurements results!

## Troponin I
Patientennr. | Vista (ng/ml) | Finecare (ng/ml) |  | Vista ng/ml | Finecare ng/ml
22646903 | 0.14 | 0.1
22648125 | 0.01 | 0.1
22647941 | 10.3 | 9.79 | AMR: | 0,015-40 | 0,1-50
22648267 | 3.87 | 1.41
22648081 | 0.042 | 0.1 | clinical decision limit: | European Society of Cardiology
22649460 | 0.127 | 0.1 |  | Cut-off 0,045
22649462 | 0.245 | 0.1
22649267 | 0.01 | 0.1 |  | American Heart Organisation
22650475 | 0.01 | 0.1 |  | values higher than 0,1 ng/ml high risk for myocardial infarction
22650444 | 0.64 | 0.1
22650464 | 0.131 | 0.1 | Problem: |  | measurement range starts only at 0,1 ng/ml
22650917 | 15.6 | 9.67
22643116 | 1.43 | 0.31
22651214 | 1.14 | 1
22652087 | 0.026 | 0.1
22652070 | 1.2 | 0.1
22652131 | 0.422 | 0.1
22653116 | 0.01 | 0.1
22653078 | 1.45 | 0.58
22653099 | 33.3 | 24.35
22653146 | 0.01 | 0.1
22653106 | 7 | 1.6
22652837 | 1.8 | 0.95
22653611 | 0.01 | 0.1
22653612 | 0.05 | 0.1
652311 | 81.3 | 28.44
22654850 | 0.023 | 0.1
22654866 | 0.01 | 0.1
22655854 | 0.01 | 0.2
22655611 | 0.91 | 0.1
22652321 | 1.12 | 6.57
22642269 | 0.071 | 0.1
22656718 | 0.01 | 0.1
22656659 | 12.6 | 4.81
22656680 | 0.247 | 0.1
22660327 | 13.7 | 8.13
22660491 | 0.026 | 0.1
22659878 | 0.01 | 0.1
22660958 | 3.69 | 1.62
22661866 | 10.5 | 6.52
22661814 | 0.023 | 0.1
22661536 | 0.098 | 0.19
22661844 | 0.063 | 0.1
22661903 | 0.315 | 0.1
22664799 | 0.106 | 0.1
22665207 | 2.76 | 1.53
22665251 | 0.155 | 0.1
22667226 | 18.3 | 5.17